## Slide 1
John A. Hole
G. D. Beskardes Q. Wu T. W. Rasmussen
M. C. Chapman K. K. Davenport
Virginia Tech
L. D. Brown D. A. Quiros R. M. Russo A. C. Stanciu
Cornell University University of Florida
Quantitative Comparison of Real Data
For Dense Geophone Arrays and Sparser Broadband Stations
Inferences for Instrument Choice
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
1 broadband
many geophones

## Slide 2
Beskardes et al., accepted
Beskardes et al., 2018 Davenport et al., 2015
Aftershock sequence of the 2011 Virginia earthquake
derived from
the dense AIDA array and backprojection
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
Rasmussen et al., in progress
Davenport et al., 2017 Stanciu et al., 2016
Comparison of
4.5 Hz geophones to broadband seismometers
from a real deployment

## Slide 3
large traditional aftershock array
36 stations
@ 2-10 km spacing
2011 Mw 5.7 Mineral, Virginia, Earthquake
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
Washington Monument
Louisa county
nuclear power plant
Davenport et al., 2015; Beskardes et al., 2018

## Slide 4
148 1-component stations
@ 200-400 m spacing
in two 6-day phases
30 3-component stations
@ 2 km spacing
3 days
AIDA
(Aftershock Imaging with Dense Arrays)
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
Davenport et al., 2015; Beskardes et al., accepted

## Slide 5
AIDA
(Aftershock Imaging with Dense Arrays)
1990’s electronics
limited power
batteries 3 days
limited memory
100 samples/second
data 6 days
 12 days deployment
1990’s “nodes”
4.5 Hz vertical component geophone
3 nodes with 3-component geophone
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
148 stations
6 people, 3 hours
7-10 stations per hour per person

## Slide 6
magnitude 2.5, 5.5 km depth
all aftershocks: signal up to 50 Hz Nyquist
Data
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
first motions
focal plane
Beskardes et al., 2018

## Slide 7
Data
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
magnitude –1.2
1.5 km depth
AIDA
400 m spacing
traditional
network
2-10 km spacing
not detected
Beskardes et al., accepted

## Slide 8
Tomography
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
>300 event tomography
spatial resolution ~1 km
limited mainly by line spacing
P & S velocity is ~constant (to 0.05 km/s)
to 16 ms (P), 60 ms (S)
VP 6.2 km/s, VS 3.6 km/s, VP/VS = 1.72
 felsic crystalline rocks within an accreted Piedmont terrane
Davenport et al., 2015
resolution test

## Slide 9
M2.5, 5.5 km depth backproject raw seismograms 1-45 Hz
Backprojection
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
spatial resolution: 75 m horizontal, 225 m vertical
detected & located events with S/N ≥ 0.6
Beskardes et al., 2018

## Slide 10
bad noise bursts (cars)
Backprojection
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
robust autodetection: backproject the kurtosis of the waveforms
then improve resolution: backproject raw data with diversity stack
M –0.7
Beskardes et al., 2018

## Slide 11
calibrate backprojection energy with
traditional event magnitude of larger events
Backprojection Magnitude
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
Beskardes et al., accepted
magnitudes from Wu et al., 2015

## Slide 12
# events:
AIDA backproj.: 1673
traditional: 813
trad. double-diff.: 494
completeness:
AIDA backproj: M –0.5
traditional: M –0.1
trad. double-diff.: M +0.2
Catalog
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
Beskardes et al., accepted
traditional from Wu et al., 2015

## Slide 13
# events:
AIDA backproj.: 1673
traditional: 813
trad. double-diff.: 494
completeness:
AIDA backproj: M –0.5
traditional: M –0.1
trad. double-diff.: M +0.2
Catalog
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
Beskardes et al., accepted
traditional from Wu et al., 2015
b-value is low (0.71)
(high stress)

## Slide 14
new result: b-value <3 km depth: 0.87
>3 km depth: 0.69
(higher stress)
Gutenberg-Richter
b-value is low (0.71)
(high stress)
Beskardes et al., accepted
traditional from Wu et al., 2015
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018

## Slide 15
AIDA backprojection
Hypocenters
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
traditional network
same major patterns
newly discovered
Beskardes et al., accepted
traditional from Wu et al., 2015
observed,
but not located
by AIDA
(outside array)

## Slide 16
Hypocenters
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
newly
discovered
AIDA backprojection
traditional network
Beskardes et al., accepted
traditional from Wu et al., 2015

## Slide 17
Hypocenters
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
newly
discovered:
not
a single
fault
 multiple
small
faults
AIDA backprojection
traditional network
Beskardes et al., accepted
traditional from Wu et al., 2015

## Slide 18
for the cost of ~3 traditional stations…
3D velocity model
2.1 times as many events detected & located
3.3 times as many events located with <200 m accuracy
newly discovered
- statistics
  - event clusters
- internal structure of main zone (small faults)
AIDA Backprojection
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018

## Slide 19
for the cost of ~3 traditional stations…
3D velocity model
2.1 times as many events detected & located
3.3 times as many events located with <200 m accuracy
newly discovered
- statistics
  - event clusters
- internal structure of main zone (small faults)
AIDA is not “Large N”
for same magnitude event in CA
more complicated velocity structure  denser lines
larger rupture  larger map area
 1000+ stations… but this is done routinely with geophones
AIDA Backprojection
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018

## Slide 20
Beskardes et al., accepted
Beskardes et al., 2018 Davenport et al., 2015
Aftershock sequence of the 2011 Virginia earthquake
derived from
the dense AIDA array and backprojection
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
Rasmussen et al., in progress
Davenport et al., 2017 Stanciu et al., 2016
Comparison of
4.5 Hz geophones to broadband seismometers
from a real deployment

## Slide 21
broadband: 15 km spacing inline map: 30 km
controlled source: 0.2 km spacing shots: 50 km
& structural geology, petrology, geochemistry
IDOR: EarthScope Idaho-Oregon 2011-2013
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
Stanciu et al., 2016; Davenport et al., 2017

## Slide 22
15 km: 1 broadband, 77 geophones
broadband station 0.7 km off geophone line
IDOR Stations
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
T. Rasmussen

## Slide 23
IDOR Seismometers
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
Guralp CMG-3T
Reftek RT125
Reftek RT130
4.5 Hz
Geospace GS-11

## Slide 24
1 half day (plus drive)
2 novices
50 geophone nodes
roadside, direct buried 0.05 m
IDOR Deployment
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
1 half day (plus drive)
3 experienced people
1 broadband station, solar
quiet site, direct buried 2 m
10 nodes per backpack

## Slide 25
M7.7, 625 km depth, D=62˚
Teleseismic Event
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
T. Rasmussen
0.03 to 20 Hz (0.05 to 32 s)
broadband
geophone
P
PcP
pP
S
sP
corrected
instrument
response
filter
error

## Slide 26
M7.7, 625 km depth, D=62˚
Teleseismic Event
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
T. Rasmussen
broadband
geophone
P
PcP
pP
S
0.03 to 0.12 Hz (8 to 32 s)
sP
2 decades below the geophone corner frequency
corrected
instrument
response
filter
error

## Slide 27
M7.7, 625 km depth, D=62˚
Teleseismic Event
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
T. Rasmussen
broadband
geophone
P
PcP
pP
0.03 to 0.12 Hz (8 to 32 s)
sP
2 decades below the geophone corner frequency
corrected
instrument
response
filter
error

## Slide 28
M7.7, 625 km depth, D=62˚
Teleseismic Event
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
T. Rasmussen
broadband
geophone
600 s window (P to S arrivals)
2 decades below the geophone corner frequency
corrected
instrument
response
corner
match

## Slide 29
Low Ambient Noise
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
T. Rasmussen
0.3 to 1.2 Hz (0.8 to 3.2 s)
broadband
geophone
1 decade below the geophone corner frequency
corrected
instrument
response

## Slide 30
Low Ambient Noise
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
T. Rasmussen
600 s window
broadband
geophone
1 decade below the geophone corner frequency
corrected
instrument
response
corner
match
trace
match
spectra

## Slide 31
M7.7, 625 km depth, D=62˚
Teleseismic Event
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
T. Rasmussen
77 geophones
P
pP
S
0.03 to 0.12 Hz (8 to 32 s)
600 s
sP
 stack to improve S/N
2 decades below the geophone corner frequency
corrected
instrument
response

## Slide 32
M7.7, 625 km depth, D=62˚
Teleseismic Event
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
T. Rasmussen
77 geophones
P
0.5 to 2 Hz (0.5 to 2 s)
20 s
13 km/s
beam-form to measure apparent velocity
and backazimuth
2 decades below the geophone corner frequency
corrected
instrument
response

## Slide 33
M7.7, 625 km depth, D=62˚
Teleseismic Event
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
T. Rasmussen
77 geophones
P
0.03 to 0.12 Hz (8 to 32 s)
PcP
pP
sP
broadband
geophone
2 decades below the geophone corner frequency
corrected
instrument
response

## Slide 34
M7.7, 625 km depth, D=62˚
Teleseismic Event
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
T. Rasmussen
broadband
geophone
600 s window (P to S arrivals)
77 geophones
2 decades below the geophone corner frequency
corrected
instrument
response
corner
match
unstacked
match
stacked

## Slide 35
Low Ambient Noise
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
T. Rasmussen
600 s window
broadband
geophone
1 decade below the geophone corner frequency
corrected
instrument
response
corner
match
trace
in progress: beamform noise
at all velocities & backazimuths, then stack
expected result: good to microseismic peak
expected
match
stacked

## Slide 36
2 orders of magnitude less expensive
Cost
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
1 broadband
many geophones

## Slide 37
zero power
Power
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
1 broadband
many geophones
solar power required
SWaP
power electronics only
next year’s tech:
 1 year, no solar
X

## Slide 38
Field Effort (Field Cost)
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018
2 orders of magnitude easier to deploy
10 nodes per backpack
1 half day (plus drive)
2 novices
50 geophone nodes
roadside, buried 0.05 m
1 half day (plus drive)
3 experienced people
1 broadband station, solar
quiet site, buried 2 m
SWaP
XX

## Slide 39
geophones record high-quality signal far below the corner frequency
SWaP! see also: Rodgers, 1994; Chapman 2009
to avoid spatial aliasing; to record the full wavefield
use low-frequency geophones at >0.15 Hz (<7 s)
better: 2 Hz geophones
2 orders of magnitude more stations at same cost
use broadband seismometers at <0.15 Hz (>7 s)
embedded deployments (full wavefield 10’s of s to 10’s of Hz)
arrays of arrays
applications that could benefit:
5 Hz to 5 s is spatially aliased with powered seismometers
earthquake hazards
volcanoes
source studies (nuclear)
 104 low-frequency geophone nodes
Instrument Recommendations
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018

## Slide 40
AIDA:
Davenport, K. K., Hole, J. A., Quiros, D. A., Brown, L. D., Chapman, M. C., Han, L., and Mooney, W. D., (2015).
Aftershock imaging using a dense seismometer array (AIDA) after the 2011 Mineral, Virginia, earthquake.
In Horton, J. W., Jr., Chapman, M. C., and Green, R. A., eds., The 2011 Mineral, Virginia, Earthquake, and
Its Significance for Seismic Hazards in Eastern North America: Geol. Soc. Amer. Special Paper 509,
273-283. doi:10.1130/2015.2509(15) .
Wu, Q., Chapman, and Beale, J. N., (2015). The aftershock sequence of the 2011 Mineral, Virginia, earthquake:
Temporal and spatial distribution, focal mechanisms, regional stress, and the role of Coulomb stress transfer.
Bull. Seismol. Soc. Amer., 105, 2521-2537.
Beskardes, G. D., Hole, J. A., Wang, K., Michaelides, M., Wu, Q., Chapman, M. C., Davenport, K. K., Brown, L. D.,
and Quiros, D. A., (2018). A comparison of earthquake backprojection imaging methods for
dense local arrays. Geophy. J. Int., 212, 1986-2002. doi:10.1093/gji/ggx520 .
Beskardes, G. D., Wu, Q., Hole, J. A., Chapman, M. C., Davenport, K. K., Brown, L. D., and Quiros, D. A., (accepted).
Aftershock sequence of the 2011 Virginia earthquake derived from the dense AIDA array and backprojection.
Bull. Seismol. Soc. Amer., accepted.
IDOR:
Stanciu, A. C., Russo, R. M., Mocanu, V. I., Bremner, P. M., Hongsresawat, S., Torpey, M. E., VanDecar, J. C., Foster,
D. A., and Hole, J. A., (2016). Crustal structure beneath the Blue Mountains terranes and cratonic
North America, Eastern Oregon and Idaho, from teleseismic receiver functions. J. Geophys. Res.,
121, 5049-5067, doi:10.1002/2016JB012989 .
Davenport, K.K., Hole, J. A., Tikoff, B., Russo, R. M., and Harder, S. H., (2017). A strong contrast in crustal architecture
from accreted terranes to craton, constrained by controlled-source seismic data in Idaho and eastern Oregon.
Lithosphere, 9, 325-340. doi:10.1130/L553.1 .
Rasmussen, T., W., (in progress). Comparison of 4.5 Hz geophones to broadband seismometers from a real deployment.
M.S. Thesis, Virginia Tech. [intention is to submit to SRL or BSSA in early 2019; please respect work in progress]
Other work documenting the geophone response below the corner frequency:
Rodgers, P. W., (1994). Self-noise spectra for 34 common electromagnetic seismometer/preamplifier pairs.
Bull. Seismol. Soc. Amer., 84, 222-228.
Chapman, M., (2009). A comparison of short-period and broadband seismograph systems in the context of the
seismology of the eastern United States. Seismol. Res. Lett., 80, 1019-1034.
References
John Hole, IRIS Seismic Instrumentation Technology Symposium, Tucson AZ, Oct. 2018